## Compatiblity Matrix _26
| Disinfectants Matrix | [5 empty] | System | System | System | System | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | CW, TEE | CW, TEE | CW, TEE | CW, TEE | CW, TEE | CW, TEE | CW, TEE
| No. | Classification | Name | Type | Disinfectanct Level | Manufacture | Body Surface | Control Panel | Control Panel (HERA, RS), '23.2~ | Monitor | CV1-8A | CV1-8AD | CMV1-10 | 3DC2-6 | 3D2-6 | 3D2-6ET | V2-6 | 3D4-8 | 3D4-8ET | V4-8 | VE4-8 | VN4-8 | 3D4-9 | 3D4-9ES | V5-9 | EV3-10B | LV3-14A | V6-12 | EV2-10A | EV2-12 | C1-4EC | CS1-4 | SC1-6 | CA1-7A | CA1-7AD | C2-4/20 | C2-4ES | C2-5 | CA2-6BM | C2-6 | C2-6IC | CN2-8 | C2-8 | C2-8ED | CA2-8A | CA2-8AD | CF2-8 | CA2-9A | CA2-9AD | C3-7EP | CA3-10A | C4-9/10ED | CN4-9 | C5-8 | CF4-9 | CA4-10M | CA1-7S | CA1-7SD | EC4-9 | EC4-9IS | ER4-9 | EV4-9 | EVN4-9 | VR5-9 | EA2-11B | EA2-11AR | EA2-11AV | EA2-11ARD | EA2-11AVD | miniER7 | E3-12A | LA2-9A | L4-7 | L3-8 | L3-12A | LA2-9S | LA3-14A | LA3-14AD | LA3-16A | LA3-16AD | LF5-13 | LA3-16AI | LS6-15 | LA3-22AI | LA4-18A | LA4-18AD | LA4-18B | LA4-18BD | LM2-18 | LM4-15B | LN5-12/40 | LN5-12 | LF5-12 | L5-12/50 | LS5-13 | L5-13/50 | L5-13 | L5-13IS | L7-16 | L7-16IS | LA2-14A | LA2-9R Vet | P1-4 | PA1-5A | PM1-6A | PN2-4 | P2-4AH | P2-4BA | PE2-4 | SP3-8 | PA3-8B | P3-8 | P3-8CA | PA4-12B | P4-12 | DP2B | CW2.0 | CW 4.0 | CW6.0 | DP8B | MMPT3-7 | MPT3-7
| 1 | Cleaner | Alkazyme | Liquid | Cleaner | Alkapharm | [80 empty] | ○ | ○
| 2 | Cleaner | Cidezyme | Liquid | Cleaner | Advanced Sterilization Products |  |  |  |  | ● | ● | ● | [8 empty] | ● |  |  | ● |  |  |  | ● | ● |  |  |  | ● | ● |  |  | ● | ○ |  |  | ○ | ● | ● | ● | ● | ● | ● | ● |  | ● | X | ● | ● | ● | ● | ○ | ○ | ● | ○ | ● | ● | ● | ○ | ○ | ◆ | ◆ | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ |  |  |  | ○ | ○ | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  | ○ | ● |  | ● | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ |  |  |  |  | ○ | ◆ | ◆
| 3 | Cleaner | trophon® companion cleaning wipes | Wipe | Cleaner | Nanosonics Limited | X | ○ | ○ | ○ | ● | ● | ● | [8 empty] | ● |  |  |  | ● | ● |  | ● | ● |  | ● |  | ○ | ○ |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ○ | ○ |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● | ● |  |  |  | ● | ● | ● | ●
| 4 | Cleaner | Enzol® | Liquid | Cleaner | Advanced Sterilization Products |  |  |  |  | ● | ● | ● | [8 empty] | ● |  |  | ● |  |  |  | ◆ | ● |  |  |  | ○ | ○ |  |  | ○ | ○ |  |  | ○ | ○ | ○ | ● | ● | ● | ● | ● |  | ○ |  | ○ | ● | ● | ○ | ◆ | ● | ○ | ○ | ● | ● | ● | ○ | ○ | ◆ | ◆ | ● | ● | ● |  |  |  | ○ | ● | ● | ● | ● | ● | ● | ● | ○ | ○ |  |  |  | ○ | ○ | ● |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ◆ |  | ○ | ◆ |  | ● | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ |  |  |  |  | ○ | ◆ | ◆
| 5 | Cleaner | Isopropyl alcohol 70% | Liquid | Cleaner | All Manufacture | X | X | ○ | ○ | ● | ● | ● | ○ | [7 empty] | ● |  |  |  | ○ |  |  | ● | ● |  |  | ○ | ● | ● |  |  | ● |  |  |  | ○ | ● | ● | ● | ● | ● | ● | ● |  | ● | X | ● | ○(20) | ○(20) | ● | ● | ● | ● | ○ | ● | ● | ● | ○ | ○ | ● | ● | ● | ● | ● |  | ○(15) | ○(15) | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ● | ● | ● | ○ | ○ | ● | ● | ● | ● |  | ● |  | ● | ○(15) | ○(15) | ○ | ○ | ● |  | ○ | ● | ○ | ● |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○(15) | ○(15) |  | ○(15)
| 6 | Cleaner | Klenzyme | Liquid | Cleaner | Steris |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ |  |  | ○ |  |  | ○(10) | ○(10) |  |  | ○ | ○ |  |  |  | ● | ● |  |  | ● | ○ |  |  | ○ | ● | ● | ● | ● | ● | ● | ● | ○ | ● | X | ● |  |  | ● |  |  | ● | ○ | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ |  |  |  | ○ | ○ | ○ |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  | ○ | ● |  | ● | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ |  |  |  |  | ○
| 7 | Cleaner | McKesson Alcohol Prep Pad Sterile | Wipe | Cleaner | McKesson Medical-Surgical Inc. |  |  |  |  | ○ | ○ | ○ | ○ | [7 empty] | ○ |  |  |  | ○ |  |  | ○ | ○ |  |  | ○ | ● | ● |  |  | ● |  |  |  | ○ | ● | ● | ● | ● | ● | ● | ● |  | ● | X | ● | ○(20) | ○(20) | ● | ● | ● | ● | ○ | ● | ● | ● | ○ | ○ | ● | ● | ● | ● | ● |  | ○(15) | ○(15) | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ● | ● |  | ● |  | ● | ○(15) | ○(15) | ○ | ○ | ● |  | ○ | ● | ○ | ● |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○(15) | ○(15) |  | ○(15)
| 8 | Cleaner | Metrizyme | Liquid | Cleaner | Metrex |  |  |  |  | ○ | ○ | ○ |  | X | X |  | X | X |  |  | ○ |  |  |  | ○ |  |  | ○ | ○ |  |  | ○ | ● | ● |  |  | ● | ○ |  |  | ○ | ● | ● | ● | ● | ● | ● | ● | ○ | ● | X | ● |  |  | ● |  |  | ● | ○ | ● | ● | ● | ○ | ○ | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | [5 empty] | ○ |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ○ | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  |  |  | ○ | ◆ | ◆
| 9 | Cleaner | Aniosyme DD1 | Liquid | Cleaner | Anios | [16 empty] | ○ | ○ |  | ○ | [7 empty] | ○ | ○ | [9 empty] | ○ | ○ | ○ | ○ | ○ |  | ○ |  |  | ○ | ○ | ○ | [7 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [19 empty] | ○ |  |  | ○
| 10 | Disinfectants | Protex disinfectant spray | Spray | Low | Parker Laboratories, Inc. | X | ○ | ○ | ○ | ○ | ○ | ○ |  |  |  | ○ |  |  | ○ |  | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ | ● | ● | ○ | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ | ○ | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ | ○ |  | ○ |  | ○
| 11 | Disinfectants | Anios Excel Surface Premium | Liquid | Low | Anios | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ |  |  | ○ |  |  | ○ |  | ○ | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ |  | X
| 12 | Disinfectants | Clinell Universal Santising wipes | Wipe | Low | GAMA Healthcare Ltd | X | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ●
| 13 | Disinfectants | Asepti-Wipes | Wipe | Low | Ecolab
| 14 | Disinfectants | Asepti-Wipes II | Wipe | Low | Ecolab | [27 empty] | ● | ● | [9 empty] | ● | ● | ● | ● | ● |  | ● |  |  |  |  | ● | [9 empty] | ● | ● | ● | ● | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [13 empty] | ● | [5 empty] | ● |  |  | ●
| 15 | Disinfectants | Bleach 5.25% | Liquid | Low | All Manufacture | [29 empty] | ○ | [65 empty] | ○
| 16 | Disinfectants | CaviWipes | Wipe | Low | Metrex | X | X | ○ | ○ | ● | ● | ● |  |  |  | ● |  |  |  |  | ● |  |  | ● | ● |  |  | ● | ● |  |  |  | X | X |  |  | X |  |  |  |  | X | X | ● | ● | ● | ● | ● |  | ● |  | X | ● | ● | ● | ● | ● | X |  | X | X | X | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | X | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● | ● | X | X |  | X |  | X |  |  |  |  | ● |  |  | ● |  | ●
| 17 | Disinfectants | Cleanisept Wipe forte | Wipe | Low | Dr.Schumacher GmbH | ○ | X | ○ | ○ | ● | ● | ● | ● | [7 empty] | ● |  |  | ● | ● |  |  | ● | ● |  |  |  | ● | ● |  |  | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ |  |  | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● |  |  | ● |  | ●
| 18 | Disinfectants | Cleanisept Wipes | Wipe | Low | Dr.Schumacher GmbH | X | X | ○ | ○ | ● | ● | ● | ● | [5 empty] | ● | ●(15) | ● |  |  |  | ● | ●(15) | ●(15) | ● | ● |  |  |  | ● | ● |  |  | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● |  | ●
| 19 | Disinfectants | gigasept® AF | Liquid | Low | Schulke |  |  |  |  | ○ | ○ | ○ |  |  |  | ○ |  |  |  |  | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ | ○(30) | ○(30) | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○(30) | ○(30) | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ | ○(30) | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ |  | ○ | [9 empty] | ○(30) | ○(30) |  | ○(30) |  | ○ | ○
| 20 | Disinfectants | Incidin Foam | Spray | Low | Ecolab | X | X |  | ○
| 21 | Disinfectants | MetriWipes | Wipe | Low | Metrex
| 22 | Disinfectants | Milton anti bacterial solution | Liquid | Low | Milton | ○ | ○ |  | ○ | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ | ○ |  |  | ● | ● |  | ○ | ● |  | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● |  | ● | ○ | ● |  |  | ● |  |  | ● | ○ | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | [5 empty] | ○ |  | ● | ● |  | ● |  | ● | ○ | ○ | ○ | ○ | ● |  | ○ | ● |  | ● |  | ○ | ○ |  | ○ | ○ | ○ | ○ | ○
| 23 | Disinfectants | Nuclean | Liquid | Low | Alkapharm |  |  |  |  | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ● | ● |  |  | ● |  |  |  | ○ | ● | ● | ● | ● | ● | ● | ● |  | ● | X | ● |  |  | ● |  |  | ● |  | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | X | X | [5 empty] | ○ |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 24 | Disinfectants | Sani-Cloth Active | Wipe | Low | PDI | ○ | X | ○ | ○ | ● | ● | ● |  |  |  | ● |  |  | ● |  | ● |  |  | ● | ● | ●(15) | ●(15) | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 25 | Disinfectants | Sani-Cloth AF Germicidal | Wipe | Low | PDI | [13 empty] | ● |  |  |  |  | ● | ● | [37 empty] | ● | ● | [24 empty] | ● | [17 empty] | ● | ● |  | ● | ● | ●
| 26 | Disinfectants | Sani-Cloth Bleach Germicidal | Wipe | Low | PDI | ○ | ○ | ○ | ○ | ● | ● | ● | ● | [5 empty] | ● |  | ● |  |  | ● | ● |  |  | ● | ● |  |  |  | ● | ● |  |  | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● |  |  | ● |  | ● |  | ● | ● | ● | ● | ● | ●
| 27 | Cleaner | Endozime | Liquid | Cleaner | The Ruhof Corporation | X | X |  | ○ | ○ | ○ | ○ | [8 empty] | ○ |  |  |  | ○ |  |  | ○ | ○ |  |  |  | ● | ● |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ |  |  | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ |  | ●
| 28 | Disinfectants | Sani-Cloth HB | Wipe | Low | PDI | X | ○ | ○ | ○ | ● | ● | ● |  | X | X |  | X | X | ● | ●(30) | ● |  |  | ◆ | ● |  |  | ● | ● |  | ● |  | ● | ● |  |  | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | X | ● | ●(30) | ●(30) | ● | ◆ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ●(30) | ●(30) | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ●(30) | ● | ● | ●(30) | ● | ● | ● |  |  |  |  | ◆ |  | ● | ◆ |  | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ●(30) | ●(30) |  | ●(30) | ●
| 29 | Disinfectants | Sani-Cloth Plus | Wipe | Low | PDI | X | X | ○ | ○ | ● | ● | ● | [8 empty] | ● |  |  | ● | ● |  |  | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● | ● | ● | ● |  | ● |  | ● |  |  |  |  | ● |  | ● | ● |  | ● |  | ● | ● |  | ● | ● | ● | ● | ●
| 30 | Disinfectants | Septiwipes | Wipe | Low | MEDICALIS | [13 empty] | ● | ●(15) |  |  |  |  | ● | ●(15) | ●(15) | [5 empty] | ● | ● | [9 empty] | ● | ● | ● | ● | ● |  | ● |  |  |  |  | ● | [9 empty] | ● | ● | ● | ● | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [13 empty] | ● | [5 empty] | ● |  |  | ●
| 31 | Disinfectants | Sono Ultrasound Wipes | Wipe | Low | Advanced Ultrasound Solutions Inc. | X | ○ | ○ | ○ | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  | ● |  |  | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 32 | Disinfectants | Ster-Bac Blu | Liquid | Low | Ecolab | [8 empty] | X | X |  | X | X | ○ | ○(50) | [10 empty] | ○ | [39 empty] | ○(50) | ○(50) | [16 empty] | ○(50) |  |  | ○(50) |  | ○ | [20 empty] | ○(50) | ○(50) |  | ○(50)
| 33 | Disinfectants | Super Sani-cloth | Wipe | Low | PDI | X | ○ | ○ | ○ | ● | ● | ● | [8 empty] | ● |  |  | ● | ● |  |  | ● | ● | ● |  |  | ● | ● |  |  | X |  |  |  |  | X | X | ● | ● | ● | ● | ● |  | ● |  | X | ● | ● | ● | ● | ● | X |  | X | X | X |  |  | ● | ● | ● | ● | ● | ● | ● | ● | X | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | X | X |  | X |  | ● | ● | ● | ● | ● | ● |  |  | ● |  | ● |  |  |  |  | ● | ● | ●
| 34 | Cleaner | Transeptic® Cleaning Solution | Spray | Cleaner | Parker Laboratories, Inc. | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ |  |  | ○ | ○ |  |  | ○ | ○ |  |  | ○ | ○ | ○ |  |  | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ | ○(20) | ○(20) | ○ | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○(15) | ○(15) | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  | ○ | ○ | ○ | ○(15) | ○ | ○ |  | ○ |  | ○ | ○(15) | ○(15) | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○(15) | ○(15) |  | ○(15)
| 35 | Disinfectants | T-Spray | Spray | Low | Pharmaceutical Innovations, Inc. | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  | ○ | ○ | ○(20) | ○(20) | ○ | ○ | ○ |  | ○ | ○ | ○ |  |  | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | X | ○ | ○(20) | ○(20) | ○ | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○(20) | ○(20) | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [5 empty] | ○ | ○(20) | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○(20) | ○(20) |  | ○(20)
| 36 | Disinfectants | T-SprayⅡ | Spray | Low | Pharmaceutical Innovations, Inc. | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  | ○ | ○ | ○(20) | ○(20) | ○ | ○ | ○ |  | ○ | ● | ● |  |  | ● | ○ |  |  | ○ | ● | ● | ● | ● | ● | ● | ● | ○ | ● | X | ● | ○(20) | ○(20) | ● | ● | ● | ● | ○ | ● | ● | ● | ○ | ○ | ● | ● | ● | ● | ● |  | ○(20) | ○(20) | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | [5 empty] | ○ | ○(20) | ● | ● | ○ | ● | ○ | ● | ○ | ○ | ○ | ○ | ● |  | ○ | ● | ○ | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○(20) | ○(20) |  | ○(20)
| 37 | Disinfectants | Anioxyde 1000 | Liquid | High | Anios | X | X | ○ | ○ | ● | ● | ● | ○ | [7 empty] | ● |  |  | ○ | ○ | ○ | ○ | ● | ● |  |  |  | ○ | ○ |  |  | X | ○ | X | X |  | X | X | ○ | ○ | ○ | ○ | ○ |  | ○ |  | X | ○ | ○ | ○ | ○ | ○ | X | ○ | X | X | X | ○ | ○ | ○ | ○ | ○ | ○ | X |  | ○ | ○ | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | X | ● | ● |  |  | ● | ○ | X | X |  | X |  | X |  |  | ○ | ○ | ○ |  |  | ○ |  | ○
| 38 | Disinfectants | Cidex | Liquid | High | Advanced Sterilization Products |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○(50) | ○ |  |  | ○ | ○(50) | ○ | ○ | ○ | ○ |  | ○ | ○ | ● | ● |  |  | ● | ○ |  |  | ○ | ● | ● | ● | ● | ● | ● | ● |  | ● | ○ | ● | ○(50) | ○(50) | ● | ● | ● | ● | ○ | ● | ● | ● | ○(50) | ○(50) | ● | ● | ● | ● | ● | ● | ○(50) | ○(50) | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ |  | ● | ● | ● | ● | ○ | ○(50) | ● | ● | ○(50) | ● | ○ | ● |  |  |  |  | ● |  | ○ | ● | ○ | ● |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○(50) |  |  | ◆ | ◆
| 39 | Disinfectants | Cidex OPA | Liquid | High | Advanced Sterilization Products |  |  |  |  | ○→● | ○→● | ● | ○ | ○ | ○ | ○ | ○ | ○ |  | ○(15) | ○→● | ○ | ○ | ○(15) | ◆ | ○(15) | ○(15) | ◆ | ● | ○ | ○ | ○ | ● | ● |  | ○ | ● | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ○ | ● | ○ | ● | ○(15) | ○(15) | ● |  |  | ◆ | ◆ | ◆ | ◆ | ◆ | ○(15) | ◆ | ◆ | ◆ | ● | ● | ● | ◆ | ○(15) | ○(15) | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ |  |  |  | ○ | ○ | ○→● | ○(15) | ● | ● | ○(15) | ● | ○ | ● | ○ | ○ | ○ | ○ | ● |  | ○ | ● | ○ | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○(15) | ○(15) | ○ | ○(15) | ○ | ◆ | ◆
| 40 | Disinfectants | Cidex Plus | Liquid | High | Advanced Sterilization Products | X | X | ○ | ○ | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ |  |  | ● | ○ | ○ | ○(25) | ○(25) |  |  | ● | ● | ○ | ○ | ○ | ● | ● | ○ | ○ | ● | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ○ | ● |  | ● | ○(25) | ○(25) | ● | ○ | ○ | ● | ○ | ● | ● | ● | ○(25) | ○(25) | ● | ● | ● | ● | ● | ○ | ○(25) | ○(25) | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | [5 empty] | ● | ○(25) | ● | ● | ○(25) | ● | ○ | ● | ○ | ○ | ○ | ○ | ● |  | ○ | ● | ○ | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○(25) | ○(25) | ○ | ○(25) | ○
| 41 | Disinfectants | Gigasept | Liquid | High | Schulke | [8 empty] | X | X |  | X | X | [6 empty] | ○(30) | [7 empty] | X | X | [9 empty] | X | X | X | X | X |  | X |  |  |  |  | X | [9 empty] | X | X | X | X | X |  |  |  |  | X | X | X | X | X | X | X
| 42 | Disinfectants | Gigasept FF | Liquid | High | Schulke |  |  |  |  | X | X | X |  | X | X |  | X | X | ○ | ○(30) | X |  |  |  | ○(30) | ○(30) | ○(30) | X | X | ○ |  |  | ○ | ○ |  | ○ | ○ |  | ★ | ★ |  | ○ | ○ | X | X | X | X | X |  | X |  | ○ | ○(30) | ○(30) | X |  |  | ○ | ○ | ○ | ○ | ○ | ○(30) | ○(30) | X | X | X | X | X |  | ○(30) | ○(30) | ○ | X | X | X | X | X | X | X | [7 empty] | X | ○(30) | ○ | ○ |  | ○ |  | ○ |  |  |  |  | X |  |  | X |  | ○ | [9 empty] | ○(30) | ○(30) | ○ | ○(30) |  | ◆ | ◆
| 43 | Disinfectants | Hibitane | Liquid | High | Regent Medical | [47 empty] | ○(30) | ○(30) | [8 empty] | ○(30) | ○(30) | [6 empty] | ○(30) | ○(30) | [16 empty] | ○(30) | [25 empty] | ○(30) | ○(30) |  | ○(30)
| 44 | Disinfectants | Metricide | Liquid | High | Metrex |  |  |  |  | ○ | ○ | ○ |  | X | X |  | ○ | ○ | ○ | ○(50) | ○ |  |  | ○ | ○(50) | ○(50) | ○(50) | ○ | ○ |  | ○ |  | ○ | ○ |  |  | X |  | ○ | ○ |  | X | X | ● | ● | ● | ○ | ○ |  | ● |  | X | ○(50) | ○(50) | ● |  |  | X |  | X | X | X | ○(50) | ○(50) | ● | ● | ● | ● | ● |  | ○(50) | ○(50) | X | ○ | ○ | ○ | ○ | ● | ● | ● | [5 empty] | ○ | ○ | ○ | ○(50) | X | X | ○(50) | X | ○ | ○ |  |  |  |  | ● |  | ○ | ● |  | ○ |  | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○(50) |  | ◆ | ◆
| 45 | Disinfectants | Metricide 28 | Liquid | High | Metrex | [13 empty] | ○ | [5 empty] | ○ | [7 empty] | ○ | ○ |  |  |  | ○ | [5 empty] | ● | ● | ● | ○ | ○ |  | ● |  |  |  |  | ● | [7 empty] | ○ | ○ | ● | ● | ● | ● | ● |  |  |  |  | ○ | ○ | ○ | ○ | ● | ● | ● | ○ | ○ | [11 empty] | ○ | [5 empty] | ● |  |  | ● |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ◆ | ◆
| 46 | Disinfectants | Metricide 30 | Liquid | High | Metrex |  |  |  |  | ○ | ○ | ○ | [8 empty] | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ● | ● | ● | ○ | ○ |  | ● |  | ○ |  |  | ● |  |  | ○ | ○ | ○ | ○ | ○ |  |  | ● | ● | ● | ● | ● |  |  |  | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ○ | ○ | ● | ● | ● |  |  | ○ |  | ○ | ○ |  | ○ |  | ○ | ○ | ○ |  |  | ● |  |  | ● |  | ○ |  | ○ | ○ | ○ | ○ | ○ | ○
| 47 | Disinfectants | Metricide OPA Plus | Liquid | High | Metrex | X | X | ○ | ○ | ○→● | ○→● | ● | ○ | [7 empty] | ○→● |  |  | ○ | ○ |  |  | ◆ | ● |  |  |  | ○ | ○ |  |  | ○ | ○ | ○ | ○ |  | ○ | ○ | ● | ● | ● | ○ | ○ |  | ● |  | ○ |  |  | ● | ◆ | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ◆ | ◆ | ● | ● | ● |  |  |  | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | X | X | ● | ● | ● |  |  | ○→● |  | ○ | ○ |  | ○ |  | ○ | X | X | X | X | ◆ |  |  | ● |  | ○ | [14 empty] | ◆ | ◆
| 48 | Disinfectants | Omnicide - FG2 | Liquid | High | Coventry Chemicals Limited |  |  |  |  | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | X | X |  |  | ● |  |  |  | ○ | ● | ● | ● | ● | ● | X | X |  | ● | ○ | ● |  |  | ● |  |  | ● | ○ | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  |  |  | ● | X | X | X | X | ● | ● | ● | ○ | ○ | [5 empty] | ○ |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 49 | Disinfectants | Omnicide 14 | Liquid | High | Metrex | [27 empty] | ○ | ○ | [9 empty] | ● | ● | ● | ○ | ○ |  | ● |  |  |  |  | ● | [9 empty] | ● | ● | ● | ● | ● |  |  |  |  | ○ | ○ | ○ | ○ | ● | ● | ● | [19 empty] | ● |  | ○ | ● |  |  |  | ○ | ○ |  | ○ | ○ | ○ | ○ | ○
| 50 | Disinfectants | Omnicide28 | Liquid | High | Metrex | [27 empty] | ○ | ○ |  |  |  | ○ | ○ | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ |  | ○ |  |  |  |  | ○ | [9 empty] | ○ | ○ | ○ | ○ | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [19 empty] | ○ |  | ○ | ○ |  |  |  | ○ | ○ |  | ○ | ○ | ○ | ○ | ○
| 51 | Disinfectants | PeraSafe* | Powder | High | Dupont | ○ | ○ | ○ | ○ | ● | ● | ● | ○ | X | X |  | X | X | ○ | ○(15) | ● |  |  | ○(15) | ○(15) | ○(15) | ○(15) | ● | ● |  |  |  | ○ | ○ |  |  | X | ○ | ○ | ○ |  | X | X | ● | ● | ● | ○ | ○ |  | ● |  | X | ○(15) | ○(15) | ● | ○ | ○ | X | X | X | X | X | ○(15) | ○(15) | ● | ● | ● | ● | ● |  | ○(15) | ○(15) | X | X | X | X | X | ● | ● | ● | ○ | ○ | ● | ● | ● |  |  | ● | ○(15) | X | X |  | X | ○ | X |  |  | ○ | ○ | ● |  |  | ● |  | X | [9 empty] | ○(15) | ○(15) |  | ○(15) |  | ◆ | ◆
| 52 | Disinfectants | Salvanios pH 7 | Liquid | Low | Anios
| 53 | Disinfectants | Salvanios pH10 | Liquid | Low | Anios | [80 empty] | ○ | ○ | [14 empty] | ○ |  |  |  |  | ○ | ○ |  | ○ | ○ | ○ | ○ | ○
| 54 | Disinfectants | Sekusept Extra | Liquid | High | Ecolab | [8 empty] | X | X |  | X | X
| 55 | Cleaner | Sekusept Plus | Liquid | Cleaner | Ecolab | [18 empty] | ○(30) | ○(30)
| 56 | Disinfectants | Sporox | Liquid | High | DSHealthcare Inc. | X | X | ○ | ○ | ● | ● | ● | X | X | X |  | X | X |  |  | ● | ○ | ○ |  |  |  |  | ● | ● |  |  | ○ | ○ | ○ | ○ |  | ○ | ○ |  |  |  | ○ | ○ | X | X | X | ○ | ○ | ○ | X | X | ○ |  |  | X |  |  | ○ |  | ○ | ○ | ○ |  |  | X | X | X | X | X | ○ |  |  | ○ | ○ | ○ | ○ | ○ | X | X | X | ○ | ○ | [5 empty] | ● |  | X | ○ |  | ○ |  | X | ○ | ○ | ○ | ○ | X |  |  | X | ○ | ○ | ○ |  |  | ○ | [9 empty] | ○
| 57 | Disinfectants | SporoxⅡ | Liquid | High | DSHealthcare Inc. | X | X | ○ | ○ | ● | ● | ● | X | X | X |  | X | X |  |  | ● | ○ | ○ |  |  |  |  | ● | ● |  |  | ○ | ● | ● | ○ |  | ● | ○ |  |  |  | ● | ● | ○ | ○ | ○ | ● | ● | ○ | ○ | X | ● |  |  | ○ |  |  | ● |  | ● | ● | ● |  |  | X | X | X | X | X | ○ |  |  | ● | ● | ● | ● | ● | X | X | X | ○ | ○ | [5 empty] | ● |  | ● | ● |  | ● |  | ● | ○ | ○ | ○ | ○ | ○ |  |  | ○ | ○ | ● | ○ |  |  | ○ | [9 empty] | ○
| 58 | Disinfectants | Steranios 2% | Liquid | High | Anios | [18 empty] | ○ | ○ | [7 empty] | ○ | ○ | [9 empty] | ● | ● | ● | ○ | ○ |  | ● |  |  |  |  | ● | [7 empty] | ○ | ○ | ● | ● | ● | ● | ● |  |  |  |  | ○ | ○ | ○ | ○ | ● | ● | ● | [5 empty] | ○ | ○ |  | ○ | [10 empty] | ● |  |  | ● |  |  |  | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | [5 empty] | ◆ | ◆
| 59 | Disinfectants | Revital-Ox® RESERT® | Liquid | High | Steris |  |  |  |  | ● | ● | ● | [8 empty] | ● |  |  | ○ | ○ | ○ | ○ | ● | ● |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● |  |  | ● | ○ | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ |  | ○ | [14 empty] | ◆ | ◆
| 60 | Disinfectants | Surfaces Hautes | Spray | High | Anios | [113 empty] | ○
| 61 | Disinfectants | Tristel Duo ULT | Foam | High | Tristel Solutions Limited | X | ○ | ○ | ○ | ● | ● | ● | ○ | [7 empty] | ● | ○ | ○ | ◆ | ◆ |  |  | ◆ | ● |  |  |  | ○ | ○ |  |  | ● | ○ | ○ | ○ | X | ● | ● | ● | ● | ● | ○ | ○ |  | ● |  | ● |  |  | ● | ○ | ○ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ● | ● | ● | ◆ |  |  | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ○ | ○ | X | X | X | ● | ● | ● | ○ | X | X |  | X |  | X | ○ | ○ | ○ | ○ | ● |  |  | ● |  | X |  |  |  | ○
| 62 | Disinfectants | Tristel Sporicidal Wipe | Wipe | High | Tristel Solutions Limited | X | ○ | ○ | ○ | ● | ● | ● | ○ | [7 empty] | ● | ○ | ○ |  | ○ |  |  | ● | ● |  |  |  | ○ | ○ |  |  | ○ | ○ | ○ | ○ |  | ○ | ○ | ● | ● | ● | ○ | ○ |  | ● |  | ○ |  |  | ● |  |  | ○ | ○ | ○ | ○ | ○ | ○ |  | ● | ● | ● | ● | ● |  | ● | ● | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ○ | ○ |  | ● | ● | ● | ● | ● |  | ○ | ○ |  | ○ |  | ○ | ○ | ○ | ○ | ○ | ● |  |  | ● |  | ○ |  |  |  | ○ | [10 empty] | ○ | ○
| 63 | Disinfectants | Rely+On Virkon* | Powder | High | LANXESS Corporation | [24 empty] | ○ | [39 empty] | ○ | [25 empty] | ○ | ○ | ○ | ○
| 64 | Disinfectants | Wavicide®-01 | Liquid | High | National Therapy Products Inc. |  |  |  |  | ● | ● | ● | ○ | X | X |  | ○ | ○ | ○ | ○(50) | ● |  |  | ○(50) | ○(50) | ○(50) | ○(50) | ● | ● |  | ○ |  | ● | ● | ○ |  | ● | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ○ | ● | ○ | ● | ○(50) | ○(50) | ● |  |  | ● | ○ | ● | ● | ● | ○(50) | ○(50) | ● | ● | ● | ● | ● |  | ○(50) | ○(50) | ● | ● | ● | ● | ● | ● | ● | ● | x | x | [5 empty] | ● | ○(50) | ● | ● | ○(50) | ● | ○ | ● |  |  |  |  | ● |  | ○ | ● |  | ● | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○(50) | ○(50) |  | ○(50) | ○ | ◆ | ◆
| 65 | Disinfectants | Gigasept PAA | Liquid | High | Schulke
| 66 | Disinfection System | Antigermix S1 | UV | High | Germitec |  |  |  |  | ★ | ★ | ★ | [8 empty] | ★ |  |  | ★ |  |  |  | ★ | ★ |  |  |  | ★ | ★ |  |  | ★ |  |  |  |  | ★ | ★ | ★ | ★ | ★ | ★ | ★ |  | ★ |  | ★ | ★ | ★ | ★ |  |  | ★ |  | ★ | ★ | ★ | ★ |  | ★ | ★ | ★ | ★ | ★ | ★ |  |  | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ |  | ★ | ★ |  |  | ★ |  | ★ | ★ |  | ★ |  | ★ |  |  |  |  | ★ |  | ★ | ★ |  | ★ |  | ★ | ★ |  | ★ | ★ | ★ | ★ | ★
| 67 | Disinfection System | trophon® EPR | Vapor | High | Nanosonics Limited |  |  |  |  | ○→★ | ○→★ | ★ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○→★ | ◆ | ◆ | ◆ | ◆ | ○ | ○ | ◆ | ★ |  | ○ | ○ | ○ | ○ |  |  | ● |  | ○ | ○ |  | ● | ● | ● | ● | ● | ○ | ○ |  | ● |  | ● | ○ | ○ | ● | ◆ | ● | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ● | ● | ● | ◆ | ○ | ○ | ● | ○ | ○ | ○ | ○ | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○→★ |  | ● | ● |  | ● |  | ● |  |  | ○ | ○ | ◆ |  |  | ● | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○
| 68 | Disinfection System | trophon®2 | Vapor | High | Nanosonics Limited |  |  |  |  | ○→★ | ○→★ | ★ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○→★ | ◆ | ◆ | ◆ | ◆ | ○ | ○ | ◆ | ★ |  | ○ | ○ | ○ | ○ |  |  | ● |  | ○ | ○ |  | ● | ● | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ● | ○ | ○ | ● | ◆ | ● | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ● | ● | ● | ◆ | ○ | ○ | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○→★ |  | ○ | ○ |  | ○ |  | ○ |  |  | ○ | ○ | ◆ |  |  | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○
| 69 | Gel | Aquasonic 100 | Gel | Gel | Parker Laboratories, Inc. |  |  |  |  | ○→● | ○→● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○→● | ○ | ○ | ○ | ○ | ○ | ○ | ○→● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ○ | ○ | ○ | ● | ○ | ○ | ○ | ○ | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○→● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● |  | ○ | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○
| 70 | Gel | Clear Image | Gel | Gel | NEXT Medical Products | [86 empty] | ○
| 71 | Gel | Natural Image | Gel | Gel | Sonotech Inc. |  |  |  |  | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ○ | ○ |  |  |  | ○ | [5 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  |  |  | ○ | [9 empty] | ○ | ○ | ○ | ○ | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  | ● | ● | ● |  |  | ● | [11 empty] | ○ |  |  | ○ |  |  | ○ | [12 empty] | ○
| 72 | Gel | Scan | Gel | Gel | Parker Laboratories, Inc. |  |  |  |  | ○ | ○ | ○ |  |  |  | ○ |  |  | ○ | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○
| 73 | Gel | Sonogel | Gel | Gel | Sonogel Vertriebs GmbH | [10 empty] | ○ |  |  | ○ | ○ |  |  |  | ○ | ○ | ○ | ○ |  |  |  | ○ | [5 empty] | ● |  |  |  |  | ● | ● | [8 empty] | ● | ○ | ○ |  |  |  | ● | ● | ● | ● | ● | ○ | ○ | [6 empty] | ○ | ○ | ● | [15 empty] | ○ | ● | ● | ○ | ● |  | ● | [9 empty] | ● | [9 empty] | ○ | ○ |  | ○
| 74 | Gel | Wavelength | Gel | Gel | National Therapy Products Inc. | [86 empty] | ○
| 75 | Disinfectants | Acrylan | Liquid | Low | Antiseptica |  |  |  |  | ○ | ○ | ○ | ○ | [7 empty] | ○ |  |  |  | ○ |  |  | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ |  |  | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ |  | ○
| 76 | Disinfectants | SEPTOL LINGETTES | Wipe | Low | ACTEON | [80 empty] | X | X
| 77 | Disinfectants | Wip' Anios Excel | Wipe | Low | Anios | ○ | ○ |  | ○ | ● | ● | ● | [8 empty] | ● |  |  | ● | ● |  |  | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● |  |  | ● | ● | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 78 | Cleaner | EMpower | Liquid | Cleaner | Metrex |  |  |  |  | ○ | ○ | ○ | [8 empty] | ○ |  |  |  | ○ |  |  | ○ | ○ |  |  | ○ | ○ | ○ |  |  | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  |  | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [5 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  |  |  | ○ | ◆ | ◆
| 79 | Disinfectants | HEXANIOS G+R | Liquid | Low | Anios |  |  |  |  | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ |  |  | ○ |  |  | ○ |  | ○ | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ |  | X
| 80 | Disinfectants | Oxivir TB Wipes | Wipe | Low | Diversey | X | ○ | ○ | ○ | ● | ● | ● | [8 empty] | ● |  |  | ○ | ○ | ○ | ○ | ● | ● |  |  |  | ● | ● |  |  | ○ |  |  |  |  | ○ | ○ | ● | ● | ● | ● | ● |  | ● |  | ○ | ○ | ○ | ● | ● | ● | ● | ○ | ● | ● | ● |  | ○ | ● | ● | ● | ● | ● | ○ | ○ | ○ | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ● |  |  | ○ | ○ | ● | ○ | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ○ | ●
| 81 | Disinfectants | Bacillol 30 Foam | Spray | Low | BODE Chemie GmbH |  |  |  |  | ○ | ○ | ○ | [8 empty] | ○ |  |  |  | ○ |  |  | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ |  |  | ○ |  |  | ○ | ○ | ○ | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ |  | ○
| 82 | Disinfectants | CaviCide | Spray | Low | Metrex | X | X | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ |  |  | ○ | ○ |  |  | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ |  |  | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ | ○ | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ |  | ○
| 83 | Disinfectants | Clinell Universal Spray | Spray | Low | GAMA Healthcare Ltd | [18 empty] | ○ | ○ | [11 empty] | ○ |  |  |  |  | ○ | ○ | [8 empty] | ○ | [5 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [16 empty] | ○ | ○ |  | ○ |  | ○ | [9 empty] | ○
| 84 | Disinfectants | gigasept® FF (new) | Liquid | High | Schulke |  |  |  |  | X | X | X | [8 empty] | X | [6 empty] | X | X |  |  |  | X | X |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ |  |  | ○ |  |  | ○ | ○ | ○ | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | X |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ |  | ○ | [14 empty] | ○ | ○
| 85 | Disinfectants | Envirocide | Spray | Low | Metrex | [31 empty] | ○ |  |  |  |  | ○ | ○ | [8 empty] | ○ | [5 empty] | ○ | ○ | ○ | ○ | ○ | [10 empty] | ○ | [16 empty] | ○ | ○ |  | ○ |  | ○ | [9 empty] | ○
| 86 | Disinfectants | MetriGuard | Spray | Low | Metrex |  |  |  |  | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ | [7 empty] | ○ |  |  |  |  | ○ | ○ | [8 empty] | ○ | [5 empty] | ○ | ○ | ○ | ○ | ○ | [10 empty] | ○ | [14 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ | [9 empty] | ○
| 87 | Disinfectants | Chlorine 1000PPM | Liquid | High | All Manufacture | [31 empty] | ○ |  |  |  |  | ○ | ○ | [8 empty] | ○ | [5 empty] | ○ | ○ | ○ | ○ | ○ | [10 empty] | ○ | [16 empty] | ○ | ○ |  | ○ |  | ○ | [9 empty] | ○
| 88 | Disinfectants | HMC NF* | Powder | High | MDD COMPANY GmbH |  |  |  |  | ○ | ○ | ○ | [8 empty] | ○ |  |  | ○ | ○ |  |  | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ |  |  | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ | ○ | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ |  | ○
| 89 | Disinfectants | SCOP'ANIOS RC | Liquid | High | Anios |  |  |  |  | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | X | X |  |  | X |  |  |  |  | X | X | ○ | ○ | ○ | ○ | ○ |  | ○ |  | X |  |  | ○ |  |  | X |  | X | X | X |  |  | ○ | ○ | ○ | ○ | ○ |  |  |  | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ |  | X | X |  | X |  | X |  |  |  |  | ○ |  |  | ○ |  | X
| 90 | Cleaner | EmPower Foam | Liquid | Cleaner | Metrex |  |  |  |  | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  | ○ | ○ | ○ |  |  | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  |  | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ● |  | ○ | ○ | ○ | ○ | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [5 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ |  | ○
| 91 | Cleaner | Metrisponge | Liquid | Cleaner | Metrex |  |  |  |  | ○ | ○ | ○ | [8 empty] | ○ |  |  |  | ○ |  |  | ○ | ○ |  |  | ○ | ○ | ○ |  |  | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  |  | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [5 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  |  |  | ○
| 92 | Sterilization system | V-PRO | Vapor | Sterilization | Steris |  |  |  |  | ★ | ★ | ★ | [8 empty] | ★ | [6 empty] | ★ | ★ |  |  |  | ★ | ★ | [9 empty] | ★ | ★ | [19 empty] | ★ | ★ | ★ | ★ | [5 empty] | ★ | ★ | ★ | ★ |  |  |  | ★ | ◆ | ★ |  |  |  |  | ★ | [21 empty] | ★ |  |  | ★
| 93 | Disinfectants | Chlordioxid Losung | Liquid | High | vitalundfitmit100 UG | [31 empty] | ○ |  |  |  |  | ○ | ○ | [8 empty] | ○ | [5 empty] | ○ | ○ | ○ | ○ | ○ | [10 empty] | ○ | [16 empty] | ○ | ○ |  | ○ |  | ○ | [9 empty] | ○
| 94 | Cleaner | Endozime Xtreme Power | Liquid | Cleaner | The Ruhof Corporation | [31 empty] | ○ |  |  |  |  | ○ | ○ | [8 empty] | ○ | [5 empty] | ○ | ○ | ○ | ○ | ○ | [10 empty] | ○ | [16 empty] | ○ | ○ |  | ○ |  | ○ | [9 empty] | ○
| 95 | Disinfectants | Surfanios Premium | Liquid | Low | Anios | [27 empty] | ○ | ○ | [9 empty] | ○ | ○ | ○ | ○ | ○ |  | ○ |  |  |  |  | ○ | [9 empty] | ○ | ○ | ○ | ○ | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [19 empty] | ○ |  |  | ○
| 96 | Disinfectants | Virkon S* | Powder | High | LANXESS Corporation | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  | ● |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ○ | ○ | ○ | ● | ● |  | ○ |  | ● | ● | ● | ○ |  |  | ● | ● | ● | ● | ● |  |  | ○ | ○ | ○ | ○ | ○ |  | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | [7 empty] | ○ |  | ● | ● |  | ● |  | ● |  |  |  |  | ○ | ○ |  | ○ |  | ●
| 97 | Disinfectants | V-Wipes | Wipe | Low | Whiteley Corporation |  |  |  |  | ● | ● | ● | [8 empty] | ● |  |  | ○ | ○ |  |  | ● | ● | [33 empty] | ○ | ○ | [23 empty] | ● | ○
| 98 | Disinfectants | terralin PAA | Liquid | High | Schulke | X | X |  | X | [23 empty] | X | X | [9 empty] | X | X | X | X | X |  | X |  |  |  |  | X | [9 empty] | X | X | X | X | X |  |  |  |  | X | X | X | X | X | X | X
| 99 | Disinfectants | Rapicide OPA/28 | Liquid | High | Cantel Medical | [18 empty] | ○ | ○ | [7 empty] | X | X |  |  | ○ |  |  |  |  | ○ | ○ | X | X | X | X | X |  | X |  | ○ |  |  | X |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | X | X | X | X | X |  |  |  | ○ | X | X | X | X | X | X | X | [8 empty] | ○ | ○ | ○ |  | ○ |  | ○ |  |  |  |  | X |  |  | X |  | ○
| 100 | Cleaner | Tristel Pre-clean wipe | Wipe | Cleaner | Tristel Solutions Limited | ○ | ○ | ○ | ○ | ● | ● | ● | [8 empty] | ● |  |  | ○ | ○ |  |  | ● | ● |  |  |  | ● | ● |  |  | ○ |  |  |  |  | ○ | ○ | ● | ● | ● | ○ | ○ |  | ● |  | ○ | ● | ● | ● |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● |  |  |  | ○ | ○ | ○ | ○ | ○ | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ○ | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ● |  |  | ● |  | ● | [14 empty] | ○ | ○
| 101 | Disinfectants | Biguacid S | Liquid | Low | Antiseptica |  |  |  |  | ○ | ○ | ○ | [8 empty] | ○ |  |  | ○ | ○ |  |  | ○ | ○ | [7 empty] | X |  |  |  |  | X | X | [8 empty] | X | [5 empty] | X |  | X | X | X | ○ | ○ | [8 empty] | X | [14 empty] | ○ | ○ | X | X |  | X |  | X | [9 empty] | X
| 102 | Disinfectants | deconex 11 UNIVERSAL | Liquid | Low | Borer Chemie AG | [27 empty] | ○ | ○ |  |  | ★ |  |  |  |  | ★ | ★ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ★ |  |  | ○ |  |  | ★ |  | ★ | ★ | ★ |  |  | ○ | ○ | ○ | ○ | ○ |  |  |  | ★ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [9 empty] | ★ | ★ |  | ★ |  | ★ |  |  |  |  | ○ |  |  | ○ |  | ★
| 103 | Disinfectants | MEDI CLEAN | Liquid | Low | UBM Cleaning Solution INC | [27 empty] | X | X |  |  | ○ |  |  |  |  | ○ | ○ | X | X | X | X | X |  | X |  | ○ |  |  | X |  |  | ○ | ○ | ○ | ○ | ○ |  |  | X | X | X | X | X |  |  |  | ○ | X | X | X | X | ○ | ○ | ○ | [9 empty] | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ |  | ○
| 104 | Disinfectants | saFecide Sensitive wipes | Wipe | Low | Cosell | X | ○ | ○ | ○ | ● | ● | ● | [8 empty] | ● |  |  |  | ● |  |  | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  |  | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 105 | Disinfectants | SURFA' SAFE | Liquid | Low | Anios | ○ | ○ | ○ | ○ | [14 empty] | ○ | ○ | [11 empty] | ○ |  |  |  |  | ○ | ○ | [8 empty] | ○ | [5 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [15 empty] | ○ | ○ | ○ |  | ○ |  | ○ | [9 empty] | ○
| 106 | Disinfectants | Mikrobac Tissues | Wipe | Low | BODE Chemie GmbH | X | X | X | ○ | ● | ● | ● | [8 empty] | ● |  |  |  | ● |  |  | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  | ● |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 107 | Disinfectants | Meliseptol Wipes Sensitive | Wipe | Low | B. Braun Medical AG | X | ○ | ○ | ○ | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | [5 empty] | ● | ● | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 108 | Disinfectants | Mikrozid sensitive liquid | Liquid | Low | Schulke |  |  |  |  | ○ | ○ | ○ | [8 empty] | ○ |  |  |  | ○ |  |  | ○ | ○ | [7 empty] | ○ |  |  |  |  | ○ | ○ | [8 empty] | ○ | [5 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [14 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ | [9 empty] | ○
| 109 | Disinfectants | Mikrozid sensitive Wipes | Wipe | Low | Schulke | X | ○ |  | ○ | ● | ● | ● | [8 empty] | ● |  |  |  | ● |  |  | ● | ● |  |  |  | ● | ● | [9 empty] | ● | ● | ● | ● | ● |  | ● |  |  | ● | ● | ● | ● | ● | [5 empty] | ● | ● | ● | ● | ● | ● | ● |  | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | [11 empty] | ● |  |  | ● | ● | [5 empty] | ● | ● | ●
| 110 | Disinfectants | Mikrozid universal liquid | Liquid | Low | Schulke | X | ○ |  | ○ | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● | [7 empty] | ○ |  |  |  |  | ○ | ○ | [8 empty] | ○ | [5 empty] | ○ | ○ | ○ | ○ | ○ | [10 empty] | ○ | [14 empty] | ● |  | ○ | ○ |  | ○ |  | ○ | [9 empty] | ○
| 111 | Disinfectants | Mikrozid universal Wipes premium | Wipe | Low | Schulke | [27 empty] | ● | ● |  |  | X |  |  |  |  | X | X |  |  |  | ● | ● |  |  |  | X | [5 empty] | X |  | X | X | X | [10 empty] | X | ● | ● | ● | ● | [12 empty] | X | X |  | X |  | ● | [9 empty] | ●
| 112 | Disinfectants | PROTEX ULTRA DISINFECTANT WIPES | Wipe | Low | Parker Laboratories, Inc. | ○ | ○ | ○ | ○ | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 113 | Disinfectants | Descogen-I | Liquid | High | Antiseptica |  |  |  |  | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ |  |  | ○ |  |  | ○ | ○ | ○ | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ |  | ○
| 114 | Disinfectants | gigasept pearls* | Powder | High | Metrex |  |  |  |  | ○ | ○ | ○ | [8 empty] | ○ |  |  |  | ○ |  |  | ○ | ○ |  |  |  | ● | ● | [9 empty] | ● | ● | ● | ○ | ○ |  | ● |  |  | ● | ● | ● | [7 empty] | ○ | ○ | ● | ● | ● | ● | ● |  |  |  |  | ○ | ○ | ○ | ○ | ● | ● | ● | [7 empty] | ○ | [11 empty] | ● |  |  | ●
| 115 | Disinfectants | Metricide Plus 30 | Liquid | High | Metrex |  |  |  |  | ● | ● | ● | [8 empty] | ● |  |  | ○ | ○ |  |  | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [5 empty] | ● | ○ | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ● | [14 empty] | ◆ | ◆
| 116 | Disinfectants | Mikrozid AF Wipes | Wipe | Low | Schulke & Mayr
| 117 | Disinfectants | CaviWipes AF | Wipe | Low | Metrex | [18 empty] | ● | ● | [37 empty] | ● | ● | [24 empty] | ●
| 118 | Disinfectants | Protex Pro Disinfection wipe | Wipe | Low | Parker Laboratories, Inc. | X | X |  | ○ | ● | ● | ● | [8 empty] | ● |  |  | ● | ● |  |  | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● | ● | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 119 | Cleaner | Tristel Rinse Wipes | Wipe | Cleaner | Tristel Solutions Limited | ○ | ○ |  | ○ | ● | ● | ● | [8 empty] | ● |  |  | ○ | ○ |  |  | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ○ | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ● | [14 empty] | ○ | ○
| 120 | Disinfectants | Clorox Healthcare® Bleach Germicidal Wipes | Wipe | Low | Clorox | ○ | ○ | ○ | ○ | ● | ● | ● | [8 empty] | ● |  |  | ● | ● |  |  | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● |  |  | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 121 | Cleaner | Sani-Cloth Detergent Wipes | Wipe | Cleaner | PDI | X | ○ | ○ | ○ | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  | ● |  |  | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 122 | Disinfectants | SciCan Optim 33TB Surface Cleaner Disinfectant Solution | Liquid | Low | SciCan |  |  |  |  | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ |  |  | ○ |  |  | ○ | ○ | ○ | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ |  | ○
| 123 | Disinfectants | SciCan Optim 33TB Surface Cleaner Disinfectant Wipes | Wipe | Low | SciCan | X | ○ | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ● | ● |  |  | ○ |  |  |  |  | ○ | ○ | ● | ● | ● | ○ | ○ |  | ● |  | ○ |  |  | ● |  |  | ○ |  | ○ | ○ | ○ |  |  | ● | ● | ● | ● | ● |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ |  | X
| 124 | Disinfectants | VERT 2 GO ED | Liquid | Low | Wood Wyant |  |  |  |  | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ |  |  | ○ |  |  | ○ | ○ | ○ | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ |  | ○
| 125 | Disinfectants | Bacillol 30 Tissue | Wipe | Low | BODE Chemie GmbH | X | X | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ○ | ○ | ○ | ○ | ○ |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ● | ● |  | ● | ● | ● | ● | ○ |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● | ●
| 126 | Disinfectants | Descoton Extra | Liquid | High | Dr.Schumacher GmbH |  |  |  |  | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ |  |  | ○ |  |  | ○ | ○ | ○ | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ |  | ○
| 127 | Disinfectants | Sekusept Aktiv* | Powder | High | Ecolab | X | X | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ● | ● |  |  | X |  |  |  |  | X | X |  |  |  | ● | ● |  |  |  | X | [5 empty] | X |  | X | X | X | [10 empty] | X | ● | ● | ● | ● | [10 empty] | ○ |  | X | X |  | X |  | X | [9 empty] | X
| 128 | Disinfectants | Clinell Universal wipes | Wipe | Low | GAMA Healthcare Ltd | X | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ●
| 129 | Disinfectants | Medilox | Liquid | High | Soosan CMC |  |  |  |  | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ | [7 empty] | X |  |  |  |  | X | X | ○ | ○ | ○ |  |  |  | ○ |  | X |  |  | ○ |  |  | X |  | X | X | X |  |  | ○ | ○ | ○ | ○ | ○ |  |  |  | X |  |  |  |  | ○ | ○ | ○ | [7 empty] | ○ |  | X | X |  | X |  | X |  |  |  |  | ○ |  |  | ○ |  | X
| 130 | Disinfectants | Clorox Healthcare Hydrogen peroxide wipes | Wipe | Low | Clorox | X | X | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ |  |  |  |  | ● | ● | ○ | ○ |  |  |  | ○ | ○ |  |  | X |  |  |  |  | X | X | ● | ● | ● | ○ | ○ |  | ● |  | X | ○ | ○ | ● |  |  | X |  | X | X | X |  | ○ | ● | ● | ● | ● | ● | ○ | ○ | ○ | X | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● |  |  |  | ● | ● | ○ | ○ | X | X |  | X |  | X |  |  |  |  | ● |  |  | ● | ● | X
| 131 | Disinfectants | Medipal Alcohol Disinfectant Wipes | Wipe | Low | Pal International limited |  |  |  |  | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ |  |  | ○ |  |  | ○ | ○ | ○ | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ |  | ○
| 132 | Disinfectants | 3M SURFACE DISINFECTANT CLEANER WIPES | Wipe | Low | 3M | X | X | ○ | ○ | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  | ● |  |  | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 133 | Disinfectants | Medipal Disinfectant Wipes | Wipe | Low | Pal International limited | X | X | ○ | ○ | [23 empty] | ● | ● |  |  | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ○ |  |  | ● | ● | ● | ● | ● | ● | ● | ● | [5 empty] | ● | ● |  | ○ | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 134 | Disinfectants | Wip' Anios Spor'Active | Foam | High | Anios | X | ○ |  | ○ | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 135 | Disinfectants | Pera flu | Liquid | High | SH |  |  |  |  | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ○ | ○ | [9 empty] | ○ | ○ | ○ | ○ | ○ |  | ○ |  |  |  |  | ○ | [9 empty] | ○ | ○ | ○ | ○ | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ | [11 empty] | ○ |  |  | ○
| 136 | Disinfectants | SURFA'SAFE PREMIUM | Spray | Low | Anios |  |  |  |  | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ● | ● |  |  | ○ | ○ |  |  |  | ○ | ○ | ● | ● | ● | ● | ● |  | ● |  | ○ | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ |  |  | ● | ● | ● | ● | ● |  |  |  | ○ | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ● |  |  | ● |  | ○
| 137 | Disinfectants | Azomax | Wipe | Low | Synergy Health Ltd | X | ○ | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○→● | ○→● | ○→● | ○ | ○ |  | ○→● |  | ○ | ● | ● | ○→● |  |  | ○ | ○ | ○ | ○ | ○ |  | ● | ○→● | ○→● | ○→● | ○→● | ○→● |  |  |  | ○ | ○ | ○ | ○ | ○ | ○→● | ○→● | ○→● | [7 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○→● |  |  | ○→● |  | ○
| 138 | Disinfectants | Sani Cloth active | Wipe | Low | Ecolab | X | X | ○ | ○ | [15 empty] | ● | [7 empty] | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  | ● | ● | ● | ● | ○ | ● | ● | ● |  | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [9 empty] | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 139 | Disinfectants | Incidin oxywipe | Wipe | Low | Ecolab | X | ○ | ○ | ○ | ● | ● | ● | [8 empty] | ● |  |  |  |  | ● | ● | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | X | X | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● | ●
| 140 | Disinfectants | Incidin oxywipe S | Wipe | Low | Ecolab | X | ○ |  | ○ | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 141 | Disinfectants | 绿先锋消毒剂 | Liquid | High | Techgreen | [68 empty] | X | X | X | X
| 142 | Cleaner | Amity 3 step solution: Detergent | Wipe | Cleaner | Amity | ○ | ○ |  | ○ | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  | ● |  |  | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 143 | Disinfectants | Amity 3 step solution: HLD | Wipe | High | Amity | ○ | X |  | ○ | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  | ● |  |  | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 144 | Cleaner | Amity 3 step solution: Rinse | Wipe | Cleaner | Amity | ○ | ○ |  | ○ | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● |  |  |  | ● | ● |  |  |  | ● | [5 empty] | ● | ● | ● | ● | ● | [10 empty] | ● | ● | ● | ● | ● | [10 empty] | ● |  | ● | ● |  | ● |  | ● | [9 empty] | ●
| 145 | Disinfectants | Incidin oxyfoam | Foam | Low | Ecolab | [27 empty] | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | [9 empty] | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 146 | Disinfectants | Incidin oxyfoam S | Foam | Low | Ecolab | X | ○ | ○ | ○ | [23 empty] | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | [9 empty] | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 147 | Disinfectants | Dispatch Hospital Cleaner Disinfectant Towels with Bleach | Wipe | Low | Clorox | ○ | ○ | ○ | ○ | ● | ● | ● | [8 empty] | ● |  |  |  | ● |  |  | ● | ● |  |  |  | ● | ● | [9 empty] | ● | ● | ● | ● | ● |  | ● |  |  | ● | ● | ● | ● | ● | [7 empty] | ● | ● | ● | ● | ● |  | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | [11 empty] | ● |  |  | ●
| 148 | Disinfection System | Hypernova Chronos | UV | High | Germitec |  |  |  |  | ★ | ★ | ★ | [8 empty] | ★ |  |  | ★ |  |  |  | ★ | ★ |  |  |  | ★ | ★ |  |  | ★ |  |  |  |  | ★ | ★ | ★ | ★ | ★ | ★ | ★ |  | ★ |  | ★ | ★ | ★ | ★ |  |  | ★ |  | ★ | ★ | ★ | ★ |  | ★ | ★ | ★ | ★ | ★ | ★ |  |  | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ |  | ★ | ★ |  |  | ★ |  | ★ | ★ |  | ★ |  | ★ |  |  |  |  | ★ |  | ★ | ★ |  | ★ |  | ★ | ★ |  | ★ | ★ | ★ | ★ | ★
| 149 | Cleaner | Sani-Cloth 70 | Wipe | Cleaner | PDI | X | ○ | ○ | ○ | ● | ● | ● |  | ○ | ○ | ○ | ● | ● | ○ | ○ | ● |  |  | ○ | ○ | ○ | ○ | ● | ● |  | ○ |  | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● |  | ● |  | ○ | ○ | ○ | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● |  | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ |  |  | ○ | ○ | ● | ○ | ○ | ○ |  | ○ |  | ○ | ○ | ○ | ○ | ○ | ● |  | ○ | ● | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○
| 150 | Disinfectants | UMONIUM 38 NEUTRALIS TISSUES | Wipe | Low | Laboratoire Huckert’s International | X | X | X | ○ | ● | ● | ● |  |  |  |  | ● | ● | ● |  | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 151 | Disinfectants | UMONIUM 38 NEUTRALIS SPRAY | Spray | Low | Laboratoire Huckert’s International | X | X | ○ | ○ | ● | ● | ● |  |  |  |  | ● | ● | ● |  | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 152 | Gel | Opti Lube Gel | Gel | Gel | Optimum Medical | [27 empty] | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [9 empty] | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 153 | Disinfectants | Septalkan | Wipe | Low | Alkapharm | ○ | ○ |  | ○ | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 154 | Disinfectants | Tristel Trio Wipes System | Wipe | High | Tristel Solutions Limited |  |  |  |  | ○ | ○ | ○ | [8 empty] | ○ |  |  |  | ○ |  |  | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ● | ● | ● | ○ | ○ |  | ● |  | ○ |  |  | ● |  |  | ○ | ○ | ○ | ○ | ○ | ○ |  | ● | ● | ● | ● | ● |  |  |  | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | [5 empty] | ● | ● | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ● |  |  | ● |  | ○
| 155 | Disinfectants | Amity 3 step solution | Wipe | High | Amity | ○ | X |  | ○ | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● |  |  |  | ● | ● |  |  |  | ● | [5 empty] | ● | ● | ● | ● | ● | [10 empty] | ● | ● | ● | ● | ● | [10 empty] | ● |  | ● | ● |  | ● |  | ● | [9 empty] | ●
| 156 | Disinfectants | Tristel Solo Wipes | Wipe | Low | Tristel Solutions Limited | ○ | ○ | ○ | ○ | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 157 | Disinfectants | ED Wipes | Wipe | Low | Dong-A ST | ○ | ○ | ○ | ○ | [23 empty] | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● |  |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 158 | Disinfectants | U Klean | Wipe | Low | Caremate | [31 empty] | ● |  |  |  |  | ● | ● | [8 empty] | ● | [5 empty] | ● |  | ● | ● | ● | [10 empty] | ● | [16 empty] | ● | ● |  | ● |  | ● | [9 empty] | ●
| 159 | Disinfectants | Accel INTERVention Wipes | Wipe | High | Diversey | ○ | X | X | ○
| 160 | Cleaner | Clinell Detergent Wipes(Yellow) | Wipe | Cleaner | GAMA Healthcare Ltd | ○ | X | ○ | ○
| 161 | Disinfectants | Clinell Sporicidal wipes | Wipe | High | GAMA Healthcare Ltd | ○ | ○ | ○ | ○
| 162 | Disinfectants | Clorox regular (Clorox regular 1 : water 10) | Liquid | Low | Clorox | ○ | ○ | ○ | ○
| 163 | Cleaner | CN'F | Liquid | Cleaner | Neo Medical | X | ○ | ○ | ○
| 164 | Disinfectants | Distel Disinfectant Wipes | Wipe | High | Tristel Solutions Ltd. | X | X | ○ | ○
| 165 | Disinfectants | Ecolab Actichlor Tablets* | Tablet | Low | Ecolab | ○ | ○ | ○ | X
| 166 | Disinfectants | Enzysept PLUS Wipes | Wipe | Low | Wooil C&Tech | ○ | ○ | ○ | ○
| 167 | Disinfectants | Ethanol 80% | Liquid | Low | all namufacture | X | ○ | ○ | ○
| 168 | Cleaner | Incidin Liquid | Liquid | Cleaner | Ecolab | X | X | ○ | ○
| 169 | Disinfectants | Lysoform® | Liquid | High | Lysoform | X | X | ○ | ○
| 170 | Disinfectants | Accel Prevention Wipes | Wipe | High | Diversey | X | ○ | ○ | ○
| 171 | Cleaner | Meliseptol rapid | Liquid | Cleaner | B. Braun Medical AG | X | X | ○ | ○
| 172 | Disinfectants | Melsept SF | Liquid | Low | B. Braun Medical AG | X | X | ○ | ○
| 173 | Disinfectants | Oxivir® Five 16 Concentrate | Liquid | High | Diversey |  | X | ○ | ○ | [91 empty] | X
| 174 | Disinfectants | perform®* | Powder | High | Schulke | X | X | ○ | ○
| 175 | Disinfectants | Pursept AF | Solution |  | Schülke & Mayr GmbH | X | X | X | X
| 176 | Cleaner | saFecide Medical wipes | Wipe | Cleaner | Liebecos | X | ○ | ○ | ○ | ● | ● | ● | [8 empty] | ● |  |  |  | ● |  |  | ● | ● |  |  |  | ● | ● | [9 empty] | ● | ● | ● | ● | ● |  | ● |  |  | ● | ● | ● | ● | ● | [6 empty] | ● | ● | ● | ● | ● | ● |  | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | [10 empty] | ● |  |  | ●
| 177 | Disinfectants | Sani-Cloth Prime | Wipe | Low | PDI | X | X | ○ | ○
| 178 | Disinfectants | Sani-Cloth Universal wipes | Wipe | Low | PDI | X | X | ○
| 179 | Cleaner | SideKick Disinfecting Wipes | Wipe | Cleaner | Stryker Medical | X | ○ | ○ | ○
| 180 | Disinfectants | Super Sani-Cloth Plus Wipes | Wipe | Low | PDI | X | ○ | ○ | ○
| 181 | Disinfectants | ALI-FLEX RTU | Liquid | Low | Lalema | X | X | ○ | ○
| 182 | Disinfectants | Tana Apesin spray | Liquid (Spray) | Low | tana Chemie GmbH | X | X | ○ | ○
| 183 | Cleaner | terralin liquid | Liquid | Cleaner | Schülke & Mayr GmbH | X | ○ | ○ | ○
| 184 | Disinfectants | Tuffie large detergent wipes | Wipe | Low | Vernacare | X | ○ | ○ | ○
| 185 | Disinfectants | Virex® II 256 | Liquid | Low | Diversey |  | X | ○ | ○
| 186 | Disinfectants | Yuhanrox regular (Yuhanrox regular 1: water 10) | Liquid | Low | Yuhan clorox | ○ | ○ | ○ | ○
| 187 | Disinfectants | Sani-Cloth AF3 Germicidal | Wipe | Low | PDI | X | X | ○ | ○
| 188 | Disinfectants | ANIOSURF Premium NPC | Liquid | Low | Anios | ○ | ○ | ○ | ○
| 189 | Disinfectants | Avert Sporicidal disinfectant cleaner | Liquid | Low | Diversey |  | ○ | ○ | ○
| 190 | Disinfectants | AZO detergent wipe | Wipe | Low | Synergy Health Ltd | X | ○ | ○ | ○
| 191 | Cleaner | Bacillol AF | Liquid | Cleaner | BODE Chemie GmbH | X | X | ○ | X
| 192 | Disinfectants | Cleanisept Wipe maxi | Wipe | Low | Dr.Schumacher GmbH | X | X | X | ○ | ● | ● | ● | ● | [5 empty] | ● | ●(15) | ● |  |  |  | ● | ●(15) | ●(15) | ● | ● |  |  |  | ● | ● |  |  | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● |  | ●
| 193 | Disinfectants | Clinell Clorox Wipes | Wipe | Low | GAMA Healthcare Ltd
| 194 | Disinfectants | Wet Wipe Triamin Disinfection | Wipe | Low | Wetwipe | ○ | ○ |  | ○ | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  | ● |  |  | ● |  | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 195 | Cleaner | Wet Wipe Ethanol 70% Disinfection | Wipe | Cleaner | Wetwipe | X | ○ |  | ○ | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  | ● |  |  | ● |  | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 196 | Cleaner | Neodisher® MediClean forte | Liquid | Cleaner | Dr.Weigert | [114 empty] | ◆ | ◆
| 197 | Cleaner | Prolystica® 2X Concentrate Enzymatic Presoak & cleaner | Liquid | Cleaner | Steris | [114 empty] | ◆ | ◆
| 198 | Disinfectants | gigasept® PAA concentrate* | Liquid | High | Schulke | X | X |  | ○ | ★ | ★ | ★ | [8 empty] | ★ |  |  |  | ★ |  |  | ★ | ★ | [7 empty] | ★ |  |  |  |  | ★ | ★ | ★ | ★ | ★ |  |  |  | ★ |  | ★ |  |  | ★ |  |  | ★ |  | ★ | ★ | ★ |  |  | ★ | ★ | ★ | ★ | ★ | ★ |  |  | ★ |  |  |  |  | ★ | ★ | ★ | [7 empty] | ★ |  | ★ | ★ |  | ★ |  | ★ |  |  |  |  | ★ |  |  | ★ |  | ★ | [14 empty] | ◆ | ◆
| 199 | Disinfectants | TD-5® | Liquid | High | CS Medical | [114 empty] | ◆ | ◆
| 200 | Disinfectants | Bacillol 30 Sensitive Tissue | Wipe | Low | BODE Chemie GmbH | X | X |  |  | ● | ● | ● | [8 empty] | ● |  |  |  | ● | ● | ● | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● | ●
| 201 | Cleaner | BIOSPOT* | Tablet | Cleaner | HYDRACHEM | ○ | ○ |  | ○
| 202 | Disinfectants | RHEOSEPT-WD plus Wipes | Wipe | Low | Wachendorff-Chemie GmbH | X | X |  | ○ | ● | ● | ● | [8 empty] | ● |  |  |  | ● | ● | ● | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● | ●
| 203 | Disinfectants | Tristel Fuse for Stella | Liquid | High | Tristel Solutions Limited | ○ | ○ |  | ○ | ★ | ★ | ★ | [8 empty] | ★ |  |  |  | ★ |  |  | ★ | ★ | [5 empty] | ● | ● | ★ |  |  |  |  | ★ | ★ | ★ | ★ | ★ |  |  |  | ★ |  | ★ |  |  | ★ |  |  | ★ |  | ★ | ★ | ★ |  |  | ★ | ★ | ★ | ★ | ★ | ★ |  |  | ★ |  |  |  |  | ★ | ★ | ★ | [7 empty] | ★ |  | ★ | ★ |  | ★ |  | ★ |  |  |  |  | ★ |  |  | ★ |  | ★
| 204 | Gel | EcoVue | Gel | Gel | HR® Pharmaceuticals, Inc |  |  |  |  | ● | ● | ● | [8 empty] | ● |  |  |  | ● | ● | ● | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● | ● |  |  |  | ● | ● | ● | ●
| 205 | Disinfectants | Rudocid®ENDO | Wipe | Low | Medidez | X | X | ○ | X | [23 empty] | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  | ● | ● | ● | ● |  | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | [9 empty] | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 206 | Disinfectants | Velox Wipes NA | Wipe | Low | Medisept | X | X | ○ | X | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● |  |  | ● |  | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [5 empty] | ● | ● | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ●
| 207 | Disinfectants | Detro San AF Wipes | Wipe | Low | Detrox | [27 empty] | ● | ● |  |  | ● |  |  |  |  | ● | ● |  |  |  | ● | ● |  |  |  | ● |  |  |  |  | ● | ● |  | ● | ● | ● | [10 empty] | ● | ● | ● | ● | ● | [6 empty] | ● | ● | ● | ● |  |  | ● | ● |  | ● |  | ● | [9 empty] | ●
| 208 | Disinfectants | Durvet Chlorhexidine | Liquid | Cleaner | Durvet | [95 empty] | X
| 209 | Disinfectants | Nolvasan®Solution Chlorhexidine diacetate | Liquid | Cleaner | Zoetis | [95 empty] | ○
| 210 | Disinfectants | Sporicidin® disinfectant solution | Liquid | Cleaner | Contec, Inc | [95 empty] | ○
| 211 | Disinfectants | Quatricide® PV Second Generation | Liquid | Low | Pharmacal Research Laboratories | [95 empty] | X
| 212 | Disinfectants | 닥터소독티슈 | Wipe | Low | Carepharm | [27 empty] | ● | ● | [9 empty] | ● | ● | ● | ● | ● |  | ● |  |  |  |  | ● | [9 empty] | ● | ● | ● | ● | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [19 empty] | ● |  |  | ●
| 213 | Gel | Aquasonic Clear Ultrasound Gel | Gel | Gel | Parker Laboratories, Inc. |  |  |  |  | ● | ● | ● | [8 empty] | ● |  |  |  | ● |  |  | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● |  |  | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● |  | ● |  |  |  |  | ● | ● | ●
| 214 | Disinfectants | Cleanisept Wipe forte Maxi | Wipe | Low | Dr.Schumacher GmbH | ○ | X |  | ○ | ● | ● | ● | ● | [7 empty] | ● |  |  | ● | ● |  |  | ● | ● |  |  |  | ● | ● |  |  | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ |  |  | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● |  |  | ● |  | ●
| 215 | Disinfectants | Sterisept Wipes | Wipes | Low | AS Chem-Pharm
| 216 | Disinfectants | Mediwipes DM | Wipes | Low | Medilab Sp. Z.o.o
0 | 216 |  | 닥터소독티슈 | Wipe | Low | Carepharm
System
x = Not compatible(DO NOT USE) ○ = Compatible
Transducer
x = Not compatible(DO NOT USE) ○ = Compatible (From Lens to Housing) ● = Compatible (From Lens upto 50cm away from the connector) Blank = Untested.(DO NOT USE) ★ = Staining may occur on housing parts; however, the acoustic performance and image quality are not affected. ○(5) = Must not be used longer than 5 minutes / ○(10) = Must not be used longer than 10 minutes ○(15) = Must not be used longer than 15 minutes / ○(20) = Must not be used longer than 20 minutes ○(25) = Must not be used longer than 25 minutes / ○(30) = Must not be used longer than 30 minutes ○(50) = Must not be used longer than 50 minutes All transducers are under tested in IPx 7 Criteria. ◆ Compatible and Validated for disinfection method *Be sure to check the manufacturer's instructions for the dilution strength and method for using powders including Perasafe. *V-PRO products include V-PRO1, V-PRO1 Plus, V-PRO Max, V-PRO Max2, V-PRO 60 sterilizer. *Be sure to check the manufacturer's instructions before using Antigermix disinfectant to prevent discoloration.

## Ref.2 (hidden)
[10 empty] | Vivid (내자) 50 평균 117칸 / P174 = 67.2% |  |  |  | Discontinued or TBD 37 평균 27칸 / P174 = 15.5% |  |  |  |  | Vivid (외자) 23 평균 42칸 / P174 = 24.1%
1 | 110 | System |  | Body Surface | 1 | 110 | System |  | Body Surface
2 | 113 | System |  | Control Panel | 2 | 113 | System |  | Control Panel
3 | 111 | System |  | Monitor | 3 | 111 | System |  | Monitor | [15 empty] | 1G | 2G | 3G | 4G
4 | 121 | Volume |  | CV1-8A | 7 | 23 | Volume | 단종 | 3D2-6 | 4 | 121 | Volume | CV1-8A | 6 | 27 | Volume | 단종 | 3DC2-6 | 18 | 83 | Volume | 외자 | EV3-10B | Biguacid S | 시험필요 | 시험필요 | 재시험(O) | 재시험(X)
5 | 121 | Volume |  | CV1-8AD | 8 | 23 | Volume | 단종 | 3D2-6ET | 5 | 121 | Volume | CV1-8AD | 7 | 23 | Volume | 단종 | 3D2-6 | 30 | 36 | Curved | 외자 | CA2-6BM | Sani-Cloth Detergent (PDI) 신규 추가품 | 시험필요 | 시험필요 | 시험필요 | 시험필요
6 | 27 | Volume | 단종 | 3DC2-6 | 10 | 25 | Volume | 단종 | 3D4-8 | 14 | 120 | Volume | VN4-8 | 8 | 23 | Volume | 단종 | 3D2-6ET | 46 | 70 | Curved | 외자 | CF4-9
7 | 23 | Volume | 단종 | 3D2-6 | 11 | 25 | Volume | 단종 | 3D4-8ET | 21 | 120 | Volume | EV2-10A | 9 | 17 | Volume | 단종 | V2-6 | 56 | 69 | Endocavity | 외자 | EA2-11B
8 | 23 | Volume | 단종 | 3D2-6ET | 15 | 10 | Volume | 단종 | 3D4-9 | 25 | 136 | Curved | CA1-7A | 10 | 25 | Volume | 단종 | 3D4-8 | 63 | 54 | Linear | 외자 | LA2-9A
9 | 17 | Volume | 단종 | V2-6 | 16 | 10 | Volume | 단종 | 3D4-9ES | 26 | 136 | Curved | CA1-7AD | 11 | 25 | Volume | 단종 | 3D4-8ET | 64 | 54 | Linear | 외자 | L4-7
10 | 25 | Volume | 단종 | 3D4-8 | 6 | 27 | Volume | 단종 | 3DC2-6 | 29 | 133 | Curved | C2-5 | 12 | 27 | Volume | 단종 | V4-8 | 73 | 57 | Linear | 외자 | LA3-16AI
11 | 25 | Volume | 단종 | 3D4-8ET | 22 | 12 | Curved | 단종 | C1-4EC | 34 | 133 | Curved | C2-8 | 13 | 21 | Volume | 단종 | VE4-8 | 74 | 55 | Linear | 외자 | LS6-15
12 | 27 | Volume | 단종 | V4-8 | 27 | 11 | Curved | 단종 | C2-4/20 | 35 | 133 | Curved | C2-8ED | 15 | 10 | Volume | 단종 | 3D4-9 | 78 | 49 | Linear | 외자 | LA4-18B
13 | 21 | Volume | 단종 | VE4-8 | 28 | 10 | Curved | 단종 | C2-4ES | 36 | 134 | Curved | CA2-8A | 16 | 10 | Volume | 단종 | 3D4-9ES | 79 | 50 | Linear | 외자 | LA4-18BD
14 | 120 | Volume |  | VN4-8 | 29 | 133 | Curved |  | C2-5 | 37 | 134 | Curved | CA2-8AD | 17 | 53 | Volume | 단/예 | V5-9 | 96 | 29 | Phased Array | 외자 | PM1-6A
15 | 10 | Volume | 단종 | 3D4-9 | 31 | 26 | Curved | 단종 | C2-6 | 38 | 133 | Curved | CF2-8 | 81 | 54 | Linear | 단종 | LM4-15B | 99 | 36 | Phased Array | 외자 | P2-4BA
16 | 10 | Volume | 단종 | 3D4-9ES | 32 | 26 | Curved | 단종 | C2-6IC | 39 | 135 | Curved | CA2-9A | 19 | 32 | Volume | 단종 | LV3-14A | 101 | 27 | Phased Array | 외자 | SP3-8
[5 empty] | 34 | 133 | Curved |  | C2-8 | 40 | 135 | Curved | CA2-9AD | 20 | 31 | Volume | 단종 | V6-12 | 102 | 42 | Phased Array | 외자 | PA3-8B
18 | 83 | Volume |  | EV3-10B | 35 | 133 | Curved |  | C2-8ED | 42 | 133 | Curved | CA3-10A | 22 | 12 | Curved | 단종 | C1-4EC | 103 | 41 | Phased Array | 외자 | P3-8
19 | 32 | Volume | 단종 | LV3-14A | 41 | 18 | Curved | 단종 | C3-7EP | 44 | 133 | Curved | CN4-9 | 23 | 18 | Curved | 단종 | CS1-4 | 104 | 41 | Phased Array | 외자 | P3-8CA
20 | 31 | Volume | 단종 | V6-12 | 43 | 20 | Curved | 단종 | C4-9/10ED | 45 | 70 | Curved | C5-8 | 24 | 20 | Curved | 단종 | SC1-6 | 105 | 33 | Phased Array | 외자 | PA4-12B
21 | 120 | Volume |  | EV2-10A | 45 | 70 | Curved |  | C5-8 | 47 | 133 | Curved | CA4-10M | 27 | 11 | Curved | 단종 | C2-4/20 | 106 | 32 | Phased Array | 외자 | P4-12
22 | 12 | Curved | 단종 | C1-4EC | 25 | 136 | Curved |  | CA1-7A | 48 | 57 | Curved | CA1-7S | 28 | 10 | Curved | 단종 | C2-4ES | 107 | 22 | CW, TEE | 외자 | DP2B
23 | 18 | Curved | 단종 | CS1-4 | 26 | 136 | Curved |  | CA1-7AD | 49 | 58 | Curved | CA1-7SD | 31 | 26 | Curved | 단종 | C2-6 | 110 | 21 | CW, TEE | 외자 | CW6.0
24 | 20 | Curved | 단종 | SC1-6 | 48 | 57 | Curved |  | CA1-7S | 50 | 133 | Endocavity | EC4-9 | 32 | 26 | Curved | 단종 | C2-6IC | 111 | 18 | CW, TEE | 외자 | DP8B
25 | 136 | Curved |  | CA1-7A | 49 | 58 | Curved |  | CA1-7SD | 52 | 133 | Endocavity | ER4-9 | 33 | 26 | Curved | 단종 | CN2-8 | 112 | 24 | CW, TEE | 외자 | MMPT3-7
26 | 136 | Curved |  | CA1-7AD | 30 | 36 | Curved |  | CA2-6BM | 53 | 133 | Endocavity | EV4-9 | 41 | 18 | Curved | 단종 | C3-7EP | 113 | 24 | CW, TEE | 외자 | MPT3-7
27 | 11 | Curved | 단종 | C2-4/20 | 36 | 134 | Curved |  | CA2-8A | 54 | 133 | Endocavity | EVN4-9 | 43 | 20 | Curved | 단종 | C4-9/10ED
28 | 10 | Curved | 단종 | C2-4ES | 37 | 134 | Curved |  | CA2-8AD | 57 | 134 | Endocavity | EA2-11AR | 51 | 102 | Endocavity | 단종 | EC4-9IS
29 | 133 | Curved |  | C2-5 | 39 | 135 | Curved |  | CA2-9A | 58 | 134 | Endocavity | EA2-11AV | 55 | 68 | Endocavity | 단종 | VR5-9
30 | 36 | Curved |  | CA2-6BM | 40 | 135 | Curved |  | CA2-9AD | 59 | 134 | Endocavity | EA2-11ARD | 62 | 25 | Endocavity | 단/예 | E3-12A
31 | 26 | Curved | 단종 | C2-6 | 42 | 133 | Curved |  | CA3-10A | 60 | 134 | Endocavity | EA2-11AVD | 84 | 16 | Linear | 단종 | LF5-12
32 | 26 | Curved | 단종 | C2-6IC | 47 | 133 | Curved |  | CA4-10M | 61 | 133 | Endocavity | miniER7 | 86 | 19 | Linear | 단종 | LS5-13
33 | 26 | Curved | 단종 | CN2-8 | 38 | 133 | Curved |  | CF2-8 | 65 | 133 | Linear | L3-8 | 88 | 26 | Linear | 단종 | L5-13
34 | 133 | Curved |  | C2-8 | 46 | 70 | Curved |  | CF4-9 | 66 | 137 | Linear | L3-12A | 89 | 26 | Linear | 단종 | L5-13IS
35 | 133 | Curved |  | C2-8ED | 33 | 26 | Curved | 단종 | CN2-8 | 67 | 137 | Linear | LA2-9S | 90 | 29 | Linear | 단종 | L7-16
36 | 134 | Curved |  | CA2-8A | 44 | 133 | Curved |  | CN4-9 | 68 | 137 | Linear | LA3-14A | 91 | 29 | Linear | 단종 | L7-16IS
37 | 134 | Curved |  | CA2-8AD | 23 | 18 | Curved | 단종 | CS1-4 | 69 | 137 | Linear | LA3-14AD | 94 | 26 | Phased Array | 단종 | P1-4
38 | 133 | Curved |  | CF2-8 | 4 | 121 | Volume |  | CV1-8A | 70 | 133 | Linear | LA3-16A | 100 | 36 | Phased Array | 단/예 | PE2-4
39 | 135 | Curved |  | CA2-9A | 5 | 121 | Volume |  | CV1-8AD | 71 | 133 | Linear | LA3-16AD | 108 | 22 | CW, TEE | 단종 | CW2.0
40 | 135 | Curved |  | CA2-9AD | 109 | 8 | CW, TEE | 단종 | CW 4.0 | 72 | 133 | Linear | LF5-13 | 109 | 8 | CW, TEE | 단종 | CW 4.0
41 | 18 | Curved | 단종 | C3-7EP | 108 | 22 | CW, TEE | 단종 | CW2.0 | 75 | 27 | Linear | LA3-22AI
42 | 133 | Curved |  | CA3-10A | 110 | 21 | CW, TEE |  | CW6.0 | 76 | 36 | Linear | LA4-18A
43 | 20 | Curved | 단종 | C4-9/10ED | 107 | 22 | CW, TEE |  | DP2B | 77 | 36 | Linear | LA4-18AD
44 | 133 | Curved |  | CN4-9 | 111 | 18 | CW, TEE |  | DP8B | 80 | 120 | Linear | LM2-18
45 | 70 | Curved |  | C5-8 | 62 | 25 | Endocavity | 단/예 | E3-12A | 82 | 133 | Linear | LN5-12/40
46 | 70 | Curved |  | CF4-9 | 57 | 134 | Endocavity |  | EA2-11AR | 83 | 133 | Linear | LN5-12
47 | 133 | Curved |  | CA4-10M | 59 | 134 | Endocavity |  | EA2-11ARD | 85 | 133 | Linear | L5-12/50
48 | 57 | Curved |  | CA1-7S | 58 | 134 | Endocavity |  | EA2-11AV | 87 | 133 | Linear | L5-13/50
49 | 58 | Curved |  | CA1-7SD | 60 | 134 | Endocavity |  | EA2-11AVD | 92 | 131 | Linear | LA2-14A
50 | 133 | Endocavity |  | EC4-9 | 56 | 69 | Endocavity |  | EA2-11B | 93 | 9 | Linear | LA2-9R Vet
51 | 102 | Endocavity | 단종 | EC4-9IS | 50 | 133 | Endocavity |  | EC4-9 | 95 | 131 | Phased Array | PA1-5A
52 | 133 | Endocavity |  | ER4-9 | 51 | 102 | Endocavity | 단종 | EC4-9IS | 97 | 133 | Phased Array | PN2-4
53 | 133 | Endocavity |  | EV4-9 | 52 | 133 | Endocavity |  | ER4-9 | 98 | 19 | Phased Array | P2-4AH
54 | 133 | Endocavity |  | EVN4-9 | 21 | 120 | Volume |  | EV2-10A
55 | 68 | Endocavity | 단종 | VR5-9 | 18 | 83 | Volume |  | EV3-10B
56 | 69 | Endocavity |  | EA2-11B | 53 | 133 | Endocavity |  | EV4-9
57 | 134 | Endocavity |  | EA2-11AR | 54 | 133 | Endocavity |  | EVN4-9
58 | 134 | Endocavity |  | EA2-11AV | 66 | 137 | Linear |  | L3-12A
59 | 134 | Endocavity |  | EA2-11ARD | 65 | 133 | Linear |  | L3-8
60 | 134 | Endocavity |  | EA2-11AVD | 64 | 54 | Linear |  | L4-7
61 | 133 | Endocavity |  | miniER7 | 85 | 133 | Linear |  | L5-12/50
62 | 25 | Endocavity | 단/예 | E3-12A | 88 | 26 | Linear | 단종 | L5-13
63 | 54 | Linear |  | LA2-9A | 87 | 133 | Linear |  | L5-13/50
64 | 54 | Linear |  | L4-7 | 89 | 26 | Linear | 단종 | L5-13IS
65 | 133 | Linear |  | L3-8 | 90 | 29 | Linear | 단종 | L7-16
66 | 137 | Linear |  | L3-12A | 91 | 29 | Linear | 단종 | L7-16IS
67 | 137 | Linear |  | LA2-9S | 92 | 131 | Linear |  | LA2-14A
68 | 137 | Linear |  | LA3-14A | 63 | 54 | Linear |  | LA2-9A
69 | 137 | Linear |  | LA3-14AD | 93 | 9 | Linear |  | LA2-9R Vet
70 | 133 | Linear |  | LA3-16A | 67 | 137 | Linear |  | LA2-9S
71 | 133 | Linear |  | LA3-16AD | 68 | 137 | Linear |  | LA3-14A
72 | 133 | Linear |  | LF5-13 | 69 | 137 | Linear |  | LA3-14AD
73 | 57 | Linear |  | LA3-16AI | 70 | 133 | Linear |  | LA3-16A
74 | 55 | Linear |  | LS6-15 | 71 | 133 | Linear |  | LA3-16AD
75 | 27 | Linear |  | LA3-22AI | 73 | 57 | Linear |  | LA3-16AI
76 | 36 | Linear |  | LA4-18A | 75 | 27 | Linear |  | LA3-22AI
77 | 36 | Linear |  | LA4-18AD | 76 | 36 | Linear |  | LA4-18A
78 | 49 | Linear |  | LA4-18B | 77 | 36 | Linear |  | LA4-18AD
79 | 50 | Linear |  | LA4-18BD | 78 | 49 | Linear |  | LA4-18B
80 | 120 | Linear |  | LM2-18 | 79 | 50 | Linear |  | LA4-18BD
81 | 54 | Linear |  | LM4-15B | 84 | 16 | Linear | 단종 | LF5-12
82 | 133 | Linear |  | LN5-12/40 | 72 | 133 | Linear |  | LF5-13
83 | 133 | Linear |  | LN5-12 | 80 | 120 | Linear |  | LM2-18
84 | 16 | Linear | 단종 | LF5-12 | 81 | 54 | Linear | 단/예 | LM4-15B
85 | 133 | Linear |  | L5-12/50 | 83 | 133 | Linear |  | LN5-12
86 | 19 | Linear | 단종 | LS5-13 | 82 | 133 | Linear |  | LN5-12/40
87 | 133 | Linear |  | L5-13/50 | 86 | 19 | Linear | 단종 | LS5-13
88 | 26 | Linear | 단종 | L5-13 | 74 | 55 | Linear |  | LS6-15
89 | 26 | Linear | 단종 | L5-13IS | 19 | 32 | Volume | 단종 | LV3-14A
90 | 29 | Linear | 단종 | L7-16 | 61 | 133 | Endocavity |  | miniER7
91 | 29 | Linear | 단종 | L7-16IS | 112 | 24 | CW, TEE |  | MMPT3-7
92 | 131 | Linear |  | LA2-14A | 113 | 24 | CW, TEE |  | MPT3-7
93 | 9 | Linear |  | LA2-9R Vet | 94 | 26 | Phased Array | 단종 | P1-4
94 | 26 | Phased Array | 단종 | P1-4 | 98 | 19 | Phased Array |  | P2-4AH
95 | 131 | Phased Array |  | PA1-5A | 99 | 36 | Phased Array |  | P2-4BA
96 | 29 | Phased Array |  | PM1-6A | 103 | 41 | Phased Array |  | P3-8
97 | 133 | Phased Array |  | PN2-4 | 104 | 41 | Phased Array |  | P3-8CA
98 | 19 | Phased Array |  | P2-4AH | 106 | 32 | Phased Array |  | P4-12
99 | 36 | Phased Array |  | P2-4BA | 95 | 131 | Phased Array |  | PA1-5A
100 | 36 | Phased Array | 단/예 | PE2-4 | 102 | 42 | Phased Array |  | PA3-8B
101 | 27 | Phased Array |  | SP3-8 | 105 | 33 | Phased Array |  | PA4-12B
102 | 42 | Phased Array |  | PA3-8B | 100 | 36 | Phased Array | 단/예 | PE2-4
103 | 41 | Phased Array |  | P3-8 | 96 | 29 | Phased Array |  | PM1-6A
104 | 41 | Phased Array |  | P3-8CA | 97 | 133 | Phased Array |  | PN2-4
105 | 33 | Phased Array |  | PA4-12B | 24 | 20 | Curved | 단종 | SC1-6
106 | 32 | Phased Array |  | P4-12 | 101 | 27 | Phased Array |  | SP3-8
107 | 22 | CW, TEE |  | DP2B | 9 | 17 | Volume | 단종 | V2-6
108 | 22 | CW, TEE | 단종 | CW2.0 | 12 | 27 | Volume | 단종 | V4-8
109 | 8 | CW, TEE | 단종 | CW 4.0 | 17 | 53 | Volume | 단/예 | V5-9
110 | 21 | CW, TEE |  | CW6.0 | 20 | 31 | Volume | 단종 | V6-12
111 | 18 | CW, TEE |  | DP8B | 13 | 21 | Volume | 단종 | VE4-8
112 | 24 | CW, TEE |  | MMPT3-7 | 14 | 120 | Volume |  | VN4-8
113 | 24 | CW, TEE |  | MPT3-7 | 55 | 68 | Endocavity | 단종 | VR5-9